VIRCELL MYCOBACTERIUM REALTIME PCR KIT

Rapid differentiation of tuberculosis and major nontuberculous mycobacterial infections directly from respiratory specimens

Executive summary

An independent prospective multicenter evaluation of the Vircell MYCOBACTERIUM REALTIME PCR KIT has been published in ASM Microbiology Spectrum, providing clinical evidence for its use in the rapid molecular diagnosis of pulmonary mycobacterial infections:

Cano-Fernandez M, Barrado L, Bermudez-Marvalc H, Cabrera-Pineda M, Diaz-De-La-Torre I, Domingo D, Hernandez S. et al. Evaluation of clinical usefulness of a commercial multiplex PCR for diagnosing mycobacterial infections: a multicenter study. Microbiol Spectr. 2026 Jul 13:e0017526. https://pubmed.ncbi.nlm.nih.gov/42439557/

This prospective multicenter study provides strong real-world evidence for the clinical value of a new multiplex real-time PCR assay designed to detect Mycobacterium tuberculosis complex and clinically relevant nontuberculous mycobacteria directly from respiratory specimens. The study evaluated the assay for the direct detection and differentiation of Mycobacterium tuberculosis complex (MTBC), Mycobacterium avium complex (MAC) and Mycobacterium abscessus complex (MABSC) in respiratory specimens. A total of 845 samples were prospectively collected and analysed across 15 clinical microbiology laboratories serving 24 hospitals in the Madrid region. Multiplex PCR results were compared with mycobacterial culture as the reference method and with Acid-Fast Bacilli (AFB) staining.

The assay showed particularly strong performance for MTBC, with 90.3% sensitivity and 98.5% specificity, increasing to 96.3% sensitivity in AFB smear-positive specimens. Specificity exceeded 99% for MAC and MABSC, supporting the reliability of positive results for the three targeted complexes. These findings support the incorporation of the assay into routine diagnostic workflows as a rapid tool for differentiating tuberculosis from major nontuberculous mycobacterial infections directly from respiratory samples.

This differentiation has direct clinical and laboratory relevance. AFB staining can rapidly demonstrate the presence of acid-fast organisms but cannot distinguish MTBC from NTM, while culture-based identification may require several weeks. Earlier molecular differentiation can support timely tuberculosis treatment and infection-control measures or, when MAC or MABSC is detected, guide the selection of further identification and susceptibility-testing procedures.

From a diagnostic-solution perspective, this publication is particularly valuable when the MYCOBACTERIUM REALTIME PCR KIT is offered as part of Vircell’s broader mycobacterial molecular portfolio. This portfolio extends beyond direct detection from respiratory specimens and includes the MTBC SPECIES REALTIME PCR KIT for further identification of MTBC, MAC, MABSC and other nontuberculous mycobacteria from culture-positive samples. The portfolio also covers whole-process and analytical AMPLIRUN® controls to support verification and routine quality management.

Introduction

Mycobacterial respiratory infections include disease caused by Mycobacterium tuberculosis complex and a broad group of environmental organisms collectively known as nontuberculous mycobacteria (NTM).

Tuberculosis remains a major global cause of infectious morbidity and mortality. At the same time, NTM pulmonary disease is becoming increasingly relevant, particularly among older patients, individuals with structural lung disease and immunocompromised populations.

Among the most clinically significant NTM are: M. avium complex, M. abscessus complex, M. kansasii and M. xenopi

The diagnosis of NTM pulmonary disease is challenging because its clinical and radiological manifestations frequently overlap with those of tuberculosis. However, the therapeutic and public-health implications differ substantially.

Tuberculosis is generally managed using standardized first-line treatment regimens and may require immediate respiratory isolation and contact investigation. By contrast, NTM infections often require prolonged and species-specific treatment, with regimens determined by the organism involved, antimicrobial susceptibility and the patient’s underlying condition.

Rapid differentiation between MTBC and NTM is therefore essential for appropriate clinical decision-making. Conventional diagnostic methods do not always provide this distinction within a clinically useful timeframe. Acid-Fast Bacilli (AFB) staining is rapid but cannot differentiate MTBC from NTM. Culture remains fundamental but may require several weeks before definitive identification is available.

Molecular testing can shorten this diagnostic interval by detecting and differentiating clinically relevant mycobacteria directly from respiratory specimens.

Assay technology and intended positioning

The Vircell MYCOBACTERIUM REALTIME PCR KIT is a multiplex real-time PCR assay designed for the simultaneous detection of: Mycobacterium tuberculosis complex (MTBC), Mycobacterium avium complex (MAC), Mycobacterium abscessus complex (MABSC) and Mycobacterium genus.

The assay includes a human RNase P internal control to monitor the suitability of the sample and molecular process.

Within MAC, the test considers organisms belonging to the M. avium–intracellulare group, including M. avium subspecies, M. intracellulare and M. chimaera.

The assay is best positioned as a rapid complement to microscopy and culture, rather than as a standalone replacement for conventional methods.

This positioning is important because, alongside the specific detection of MTBC, MAC and MABSC, the assay includes a genus-level Mycobacterium target that can indicate the presence of other mycobacterial species not covered by those three complexes. A positive result can provide early, clinically relevant differentiation. However, a negative result does not exclude infection, particularly when the organism is present below the assay’s detection threshold, or detection is affected by low bacillary burden, sample heterogeneity, or pre-analytical or analytical factors.

In routine practice, the assay can therefore be used to accelerate early diagnostic orientation while culture continues in parallel.

Evidence from the publication

Study design

Cano-Fernandez et al. conducted a prospective multicenter study between June 2024 and July 2025. The study involved 15 microbiology laboratories processing mycobacterial cultures for 24 hospitals in the Madrid region. Respiratory specimens were collected from patients with clinical suspicion of mycobacterial infection and were processed according to each laboratory’s routine procedures.

| Study component | Details from publication |
| --- | --- |
| Study design | Prospective multicenter clinical evaluation |
| Study period | June 2024 to July 2025 |
| Participant laboratories | 15 |
| Hospital represented | 24 |
| Total respiratory specimens | 845 |
| Main sample type | Sputum (594) |
| Other sample matrices | Bronchial aspirates (146) Bronchoalveolar lavage samples (98) Bronchial lavage specimens (7) |
| Reference method | Culture |
| Additional reference | Acid-Fast Bacilli staining (AFB) |

| Study component | Details from publication |
| --- | --- |
| Molecular assay | MYCOBACTERIUM REALTIME PCR KIT |
| Extraction workflow | MagXtract® 3200 instrument Vircell Standard Extraction Plate SE001 |
| PCR instrument used | Bio-Rad CFX96 Touch |

Diagnostic workflow

Each specimen was evaluated using parallel diagnostic methods: Acid-fast bacilli staining, Mycobacterial culture and the Vircell Multiplex real-time PCR.

For molecular testing, approximately 1 mL of sample underwent thermal inactivation at 95°C for 30 minutes. A 600 µL aliquot was then processed with the MagXtract® 3200 automated system using the pre-dispensed Vircell Standard Extraction Plate Ref. SE001

The extraction procedure included proteinase K digestion and magnetic bead-based purification, producing an 80 µL eluate in approximately 20 minutes per batch. Five microlitres of extracted nucleic acid were automatically transferred to lyophilized PCR tubes before amplification.

Laboratories participating in this multicentric study were additionally queried regarding the use of other molecular assays for MTBC detection.

Microbiological findings

| Culture result | Number | AFB positive | AFB negative | Multiplex PCR with AFB positive | Multiplex PCR with AFB negative |
| --- | --- | --- | --- | --- | --- |
| MTBC | 156 | 138 | 18 | 131 | 8 |
| MAC | 120 | 79 | 41 | 60 | 12 |
| MABSC | 25 | 21 | 4 | 15 | 3 |
| Negative | 488 | 29 | 459 | 7 | 5 |

Table 1 Distribution of AFB smear and Vircell Mycobacterium  PCR results among culture-positive specimens, stratified by mycobacterial group. Cultures yielding Mycobacterium species outside the assay targets and contaminated cultures were excluded from the analysis.

The study included 845 respiratory specimens from patients with suspected mycobacterial infection: sputum (n = 594), bronchial aspirates (n = 146), bronchoalveolar lavage (n = 98), and bronchial lavage (n = 7). Median patient age was 66 years, and 50.06% were female. Culture was positive in 301 specimens (35.6%): MTBC (n = 156), MAC (n = 120), and MABSC (n = 25); 238 of these (79.1%) were also AFB-positive. AFB staining was positive but culture negative in 29 cases. PCR was positive despite negative culture and AFB staining in five cases—three MABSC and two MTBC. Overall, multiplex PCR detected mycobacterial complexes in 245 specimens (29%): MTBC (n = 147), MAC (n = 77), and MABSC (n = 21). Table 1.

Diagnostic performance

Culture was used as the reference method for calculating sensitivity, specificity, positive predictive value and negative predictive value. Table 2.

MTBC performance

The assay showed strong diagnostic performance for MTBC, with 90.26% overall sensitivity, 98.53% specificity, 95.20% positive predictive value and 96.55% negative predictive value. Notably, sensitivity increased to 96.32% among AFB smear-positive specimens.

This performance in terms of sensitivity, specificity and positive predictive value supports the use of the Vircell assay for rapid MTBC detection, particularly in patients with a higher bacillary burden. In this setting, an early molecular result may support timely treatment initiation and infection-control decisions before culture identification becomes available. Furthermore, comparison of the evaluated PCR panel with other molecular assays routinely used in the participating laboratories for MTBC detection, showed a high degree of concordance with the Vircell PCR assay, particularly among positive samples.

MAC performance

For MAC, overall sensitivity was 61.02%, increasing to 76.92% in AFB-positive specimens. Specificity was 99.37%, and the positive predictive value was 94.74%.

The lower overall sensitivity means that a negative MAC result should not be used independently to exclude infection, particularly in patients with strong clinical or radiological suspicion and a low bacillary burden. This limitation has also been reported by other publications with other NTM PCR assays and likely reflects the heterogeneity and lower bacterial burden often associated with NTM pulmonary disease. The high specificity and positive predictive value nevertheless support the clinical relevance of a positive MAC result in the evaluated population.

MABSC performance

For MABSC, sensitivity was 72.00% and specificity was 99.58%. The study included only 25 MABSC culture-positive specimens, resulting in wider confidence intervals than those observed for MTBC and MAC. The result should therefore be interpreted as supportive but less precisely estimated.

| Target | Sensitivity | Sensitivity in AFB-positive samples | Specificity | PPV | NPV |
| --- | --- | --- | --- | --- | --- |
| MTBC | 90.26% | 96.32% | 98.53% | 95.20% | 96.55% |
| MAC | 61.02% | 76.92% | 99.37% | 94.74% | 91.01% |
| MABSC | 72.00% | 71.43% | 99.58% | 85.71% | 98.76% |
| Total as reported | 89.66% | 89.36% | 9748% | 94.24% | 87.18% |

Table 2:  Diagnostic performance of Mycobacterium real-time PCR assay for detection of mycobacteria compared with culture.

A positive MABSC result may provide clinically useful early information because infections caused by the M. abscessus complex can require complex, prolonged and highly specific therapeutic approaches.

Performance according to AFB smear status

The study demonstrated that assay sensitivity was generally higher in AFB smear-positive samples, particularly for MTBC and MAC.

This relationship reflects the effect of bacillary burden on molecular detection. AFB-positive specimens usually contain a higher concentration of mycobacteria, increasing the probability that target nucleic acid will be present in the aliquot analysed by PCR.

For MTBC, sensitivity increased from 90.26% overall to 96.32% in smear-positive specimens. For MAC, sensitivity increased from 61.02% overall to 76.92% in smear-positive specimens. This result reinforces the value of combined interpretation:

- AFB staining provides rapid evidence of acid-fast organisms and an indirect estimate of bacillary burden.

- PCR provides early differentiation among the targeted complexes.

- Culture confirms viable organisms and permits broader identification and susceptibility testing.

The assay’s strongest immediate clinical contribution is therefore likely to occur when an AFB-positive respiratory specimen requires rapid differentiation between tuberculosis and NTM infection.

Why specimens are decontaminated before culture

Respiratory specimens originate from non-sterile anatomical sites and may contain large numbers of commensal bacteria and fungi.

These microorganisms generally grow faster than mycobacteria and can overgrow the culture medium, potentially preventing the recovery or recognition of the target organism. Respiratory specimens therefore undergo digestion and decontamination before culture inoculation. The objectives of this procedure are to liquefy mucus and homogenize the specimen, release mycobacteria trapped within respiratory material, reduce bacterial and fungal contamination and concentrate viable mycobacteria before inoculation.

The procedure depends partly on the relative resistance of MTBC to alkaline decontaminating reagents. However, the conditions must be carefully controlled. Insufficient decontamination can lead to bacterial or fungal overgrowth. Excessive decontamination can reduce mycobacterial viability and compromise culture sensitivity.

This effect is particularly relevant for NTM. The publication explains that some NTM, especially rapidly growing mycobacteria, may be more susceptible than MTBC to the reagents used during decontamination. Viable NTM cells may therefore be damaged or destroyed during processing, reducing their recovery in culture.

In the study, 19 cultures were contaminated with bacteria or fungi and were excluded from the performance analysis. Two of these specimens produced positive PCR results, one for MAC and one for MTBC.

This illustrates the complementary nature of culture and PCR: Culture requires viable organisms capable of growth while PCR detects target nucleic acid and is not dependent on microbial viability.

Interpretation of discrepant results

PCR-positive, culture-negative results

Seven specimens were positive for MTBC by the Vircell PCR assay but negative by culture. Among these specimens four were AFB smear positive, one came from a patient with a previous episode of tuberculosis and in four additional cases they were also positive with another molecular assay used by the participating laboratories (eg GeneXpert MTB/RIF), supporting the PCR findings obtained by the Vircell assay.

These findings support the reliability of at least some PCR-positive, culture-negative results.

Potential explanations include reduced organism viability after treatment, loss of viable organisms during decontamination or sample processing, low concentrations of viable organisms., persistence of mycobacterial DNA from a previous or treated infection and differences between the aliquots used for molecular testing and culture.

The detection of nucleic acid does not always establish active disease. Positive PCR findings should therefore be interpreted alongside the patient’s clinical presentation, radiological findings, treatment history and other microbiological results.

Two MAC PCR-positive, culture-negative results were obtained from patients with a documented history of M. intracellulare infection. This finding may also reflect persistent DNA or organisms not recovered in the culture aliquot.

Culture-positive, PCR-negative results

A negative PCR result does not exclude mycobacterial infection. Potential explanations include low bacillary burden, heterogeneous distribution of organisms within the specimen, differences between PCR and culture aliquots, PCR inhibition or inefficient nucleic acid recovery or infection caused by a species outside the targeted complexes.

In this study culture identified 37 isolates corresponding to species not specifically covered   in the PCR panel. These included M. gordonae, M. fortuitum, M. mucogenicum, M. lentiflavum, M. scrofulaceum, M. celatum, M. chelonae, M. interjectum, M. shottsii and M. xenopi. Only two of these samples were positive by AFB staining. This result reinforces the continued need for culture when the clinical suspicion of mycobacterial disease remains high.

Clinical and laboratory implications

Rapid diagnostic orientation

The principal clinical value of the MYCOBACTERIUM REALTIME PCR KIT is the early differentiation of MTBC from MAC and MABSC directly in respiratory specimens. This distinction is particularly relevant when AFB staining is positive, because microscopy confirms the presence of acid-fast organisms but does not determine whether the patient is likely to have tuberculosis or an NTM infection.

Support for early tuberculosis decisions

The high MTBC sensitivity observed in AFB smear-positive specimens (96.32%), together with 98.53% specificity, supports rapid action while culture is still pending. In an appropriate clinical context, an early MTBC-positive result may contribute to treatment initiation, respiratory isolation and other infection-control decisions. The result should remain integrated with clinical, radiological and microbiological information.

Earlier orientation for major NTM

Direct detection of MAC or MABSC can provide an early indication that the organism is a major NTM complex rather than MTBC. This may help laboratories prioritize subsequent culture identification and susceptibility-testing procedures and may support earlier specialist assessment. However, molecular detection alone does not establish that ATS/ERS/ESCMID/IDSA criteria for NTM pulmonary disease have been fulfilled.

Complementary use with microscopy and culture

The study supports an integrated workflow in which AFB staining estimates bacillary burden, multiplex PCR provides rapid differentiation among the targeted complexes, and culture confirms viable organisms, detects species outside the panel and enables susceptibility testing. PCR should therefore accelerate diagnostic orientation while culture continues in parallel, rather than replace it.

Laboratory resilience in difficult cultures

PCR can provide useful information when culture is compromised by contamination, loss of organism viability during decontamination, prior antimicrobial treatment or uneven distribution of organisms between aliquots. Conversely, culture remains essential when PCR is negative but clinical suspicion persists, especially for low-burden infection or disease caused by an NTM species outside the assay targets.

Balanced interpretation and study limitations.

The publication provides strong prospective, multicentre evidence, but the results should be interpreted within the following constraints:

- The evaluation included respiratory specimens only. Performance cannot be extrapolated directly to extrapulmonary matrices without additional validation.

- All participating laboratories were located in the Madrid region. The multicentre design improves real-world representativeness, but local prevalence and NTM epidemiology may affect predictive values and generalisability to other settings.

- Culture was used as the reference method, although culture can be affected by contamination, prior treatment, low organism viability and decontamination-related loss, particularly for some NTM. Apparent PCR-positive/culture-negative results are therefore not necessarily false-positive molecular results.

- Sensitivity was lower for MAC (61.02% overall) and MABSC (72.00%) than for MTBC, and negative results require caution in low-burden or strongly suspected NTM disease. The MABSC estimate was based on only 25 culture-positive specimens and consequently had wider confidence intervals.

- The panel directly differentiates MTBC, MAC and MABSC, but it does not identify every clinically relevant NTM. Culture detected 37 isolates belonging to species outside the targeted complexes, reinforcing the need to maintain culture and broader identification methods.

- Comparison with other rapid molecular assays was available only for a subset of samples and was not performed systematically using one uniform comparator across all centres.

- Systematic clinical and radiological correlation was not available. The study therefore evaluates microbiological performance and diagnostic usefulness, but cannot determine whether every molecular detection represented active disease, colonisation or persistent DNA after previous infection or treatment.

- The PCR results were not reported to clinicians and did not influence patient management. The study therefore did not directly measure changes in time to treatment, isolation decisions, length of stay or patient outcomes.

Conclusion

This prospective multicentre study, involving 845 respiratory specimens across 15 laboratories, provides published clinical evidence supporting the MYCOBACTERIUM REALTIME PCR KIT as a rapid molecular tool for the direct detection and differentiation of MTBC, MAC and MABSC.

The assay demonstrated its strongest performance for MTBC, with 90.3% sensitivity, increasing to 96.3% in AFB smear-positive specimens, and 98.5% specificity. Sensitivity was 61.0% for MAC, rising to 76.9% in AFB-positive samples, and 72.0% for MABSC. Specificity exceeded 99% for both NTM targets, providing high confidence in positive results. The lower sensitivity observed for NTM is consistent with the lower and more heterogeneous bacillary burden commonly associated with these infections.

By providing early differentiation between tuberculosis and major NTM infections directly from respiratory specimens, the assay can support faster and more targeted clinical, infection-control and laboratory decisions while culture continues in parallel.

Its integration with MagXtract® 3200 automated extraction and compatibility with open real-time PCR platforms provide laboratories with a standardized, flexible workflow that can be adapted to their existing molecular infrastructure and substantially reduce the time to clinically relevant results compared with culture-based identification.

Vircell as a full Mycobacterium molecular portfolio partner, not a single-test supplier

Vircell provides a comprehensive molecular portfolio for mycobacterial diagnosis, covering direct detection, species differentiation, automated extraction, quality controls and workflow connectivity detailed in Table 3.

Direct detection from respiratory specimens

The MYCOBACTERIUM REALTIME PCR KIT provides simultaneous detection of Mycobacterium genus and direct differentiation of MTBC, MAC and MABSC. It is positioned for rapid first-line molecular orientation while microscopy and culture proceed in parallel.

Further identification from positive cultures

The MTBC SPECIES REALTIME PCR KIT extends the workflow beyond initial complex-level detection by supporting further molecular identification from culture-positive material. Together, the two assays allow laboratories to connect direct detection with downstream characterisation rather than relying on an isolated test result.

Integrated automated extraction

The workflow evaluated in the publication used the MagXtract® 3200 with the pre-dispensed Vircell Standard Extraction Plate SE001. Automated magnetic-bead extraction and PCR setup support standardisation, traceability and efficient batch processing within an open real-time PCR workflow.

Quality-control support

Vircell AMPLIRUN® controls complement the molecular assays with characterised materials for verification, analytical quality control, operator training and routine monitoring. Whole-process controls can challenge extraction and amplification together, while analytical controls can support target-specific PCR checks according to the laboratory’s quality plan.

Open-platform flexibility

| The evaluated assay was amplified on a Bio-Rad CFX96 Touch, illustrating a workflow that can be implemented on an open real-time PCR platform. This gives laboratories greater flexibility than a single closed-cartridge approach and allows the mycobacterial solution to be integrated into broader molecular testing infrastructure. Together, these components position Vircell as a portfolio partner for the mycobacterial molecular pathway: rapid direct detection, downstream identification, automated extraction and quality-control support. The value proposition is therefore not limited to one multiplex Description | Reference | Class | Pack size |
| --- | --- | --- | --- |
| MYCOBACTERIUM REALTIME PCR KIT | RTPCR016-LPD | CE0123 IVDR | 96 Test |
| MYCOBACTERIUM REALTIME PCR KIT | RTPCR016-LPD | CE0123 IVDR | 24 Test |
| MTBC SPECIES REALTIME PCR KIT | RTPCR022-LP-R | RUO | 24 Test |
| MagXtract® 3200 | MAGXTRACT3200 | CE |  |
| STANDARD EXTRACTION PLATE | SE001 | CE IVDR | 96 Test |
| STANDARD EXTRACTION TUBE | SE002 | CE IVDR | 96 Test |
| VIRCOM PCR | VPCOM | CE |  |
| AMPLIRUN® TOTAL MDR-TB VERIFICATION & CONTROL PANEL (SPUTUM) | MBTC027 | CE | 5x2 vials |
| AMPLIRUN® TOTAL MTB CONTROL (SPUTUM) | MBTC013 | CE | 10 vials |
| AMPLIRUN® TOTAL MTB INH RESISTANT CONTROL (SPUTUM) | MBTC015 | CE | 2x5 vials |
| AMPLIRUN® TOTAL MTB RIF RESISTANT CONTROL (SPUTUM) | MBTC014 | CE | 2x5 vials |
| AMPLIRUN® MYCOBACTERIUM AVIUM DNA CONTROL | MBC086-R | RUO | 1 vial |
| AMPLIRUN® MYCOBACTERIUM INTRACELLULARE DNA CONTROL | MBC087-R | RUO | 1 vial |
| AMPLIRUN® MYCOBACTERIUM KANSASII DNA CONTROL | MBC095-R | RUO | 1 vial |
| AMPLIRUN® MYCOBACTERIUM TUBERCULOSIS DNA CONTROL | MBC034 | CE0123 IVDR | 1 vial |
| AMPLIRUN® MYCOBACTERIUM ULCERANS DNA CONTROL | MBC094-R | RUO | 1 vial |

Table 3: Vircell portfolio for Mycobacterium molecular diagnosis

assay, but to a connected workflow that laboratories can adapt to their diagnostic algorithms, throughput and quality requirements.

Middleware communication system

VIRCOM Molecular supports end-to-end traceability and standardized data management throughout the molecular workflow. It enables full integration with the laboratory information system (LIS) and provides automatic, straightforward interpretation of results, including amplification curves and Ct values. Human error is reduced through scanner-based data entry, automatic creation of run templates with the corresponding controls, recording of reagent lots and expiry dates, colour-coded identification of positive and negative results, and a guided result-validation wizard.

Recommended message for end users

1. Detect and differentiate early.

Use the MYCOBACTERIUM REALTIME PCR KIT to obtain rapid molecular differentiation of MTBC, MAC and MABSC directly from respiratory specimens.

2. Interpret AFB and PCR together.

AFB staining indicates the presence and approximate burden of acid-fast organisms; PCR adds early differentiation among the targeted complexes.

3. Continue culture in parallel.

Maintain culture for recovery of viable organisms, detection of species outside the PCR panel, definitive identification and antimicrobial susceptibility testing.

4. Act on positive results in context.

A positive MTBC result may support timely tuberculosis treatment and infection-control decisions. Positive MAC or MABSC results can guide further NTM investigation, but must be correlated with clinical, radiological and microbiological criteria.

5. Do not use a negative PCR result alone to exclude disease.

Low bacillary burden, sample heterogeneity, inhibition, suboptimal nucleic-acid recovery or an NTM outside the targeted complexes may produce a negative result despite ongoing infection.

6. Build a complete workflow.

Combine Vircell direct-detection and culture-identification assays with MagXtract® automated extraction and AMPLIRUN® quality controls for a standardised, traceable and flexible mycobacterial molecular solution.

References

1. Cano-Fernandez M, Barrado L, Bermudez-Marvalc H, Cabrera-Pineda M, Diaz-De-La-Torre I, Domingo D, Hernandez S, et al. Evaluation of clinical usefulness of a commercial multiplex PCR for diagnosing mycobacterial infections: a multicenter study. Microbiol Spectr. 2026; e00175-26. doi:10.1128/spectrum.00175-26.

2. Aguilera Franco M, Sampedro Padilla A, Gutierrez-Bautista JF, Gonzalez Martinez A, Cobo F, Gutierrez-Fernandez J, et al. Evaluation of the Mycobacterium RealTime Kit Vircell assay for detecting Mycobacterium species in clinical specimens. Pathogens. 2025;14:429. doi:10.3390/pathogens14050429.

3. Johansen MD, Herrmann J-L, Kremer L. Non-tuberculous mycobacteria and the rise of Mycobacterium abscessus. Nat Rev Microbiol. 2020;18:392–407. doi:10.1038/s41579-020-0331-1.

4. Ratnatunga CN, Lutzky VP, Kupz A, Doolan DL, Reid DW, Field M, et al. The rise of non-tuberculosis mycobacterial lung disease. Front Immunol. 2020;11:303. doi:10.3389/fimmu.2020.00303.

5. Daley CL, Iaccarino JM, Lange C, Cambau E, Wallace RJ Jr, Andrejak C, et al. Treatment of nontuberculous mycobacterial pulmonary disease: an official ATS/ERS/ESCMID/IDSA clinical practice guideline. Eur Respir J. 2020;56:2000535. doi:10.1183/13993003.00535-2020.

6. Kim J-U, Ryu D-S, Cha C-H, Park S-H. Paradigm for diagnosing mycobacterial disease: direct detection and differentiation of Mycobacterium tuberculosis complex and non-tuberculous mycobacteria in clinical specimens using multiplex real-time PCR. J Clin Pathol. 2018;71:774–780. doi:10.1136/jclinpath-2017-204945.

7. Broncano-Lavado A, Barrado L, Lopez-Roa P, Cacho J, Domingo D, Hernandez S, et al. Clinical evaluation of nontuberculous mycobacteria (NTM) elite agar, a new medium for the isolation of NTM: a multicenter study. J Clin Microbiol. 2023;61:e00036-23. doi:10.1128/jcm.00036-23.

8. Martin I, Pfyffer GE, Parrish N, Carroll KC, Pfaller MA, Landry ML, et al. Mycobacterium: general characteristics, laboratory detection, and staining procedures. In: Manual of Clinical Microbiology. 12th ed. Washington, DC: ASM Press; 2019. p. 558–575.

9. Buijtels PCAM, Petit PLC. Comparison of NaOH-N-acetyl cysteine and sulfuric acid decontamination methods for recovery of mycobacteria from clinical specimens. J Microbiol Methods. 2005;62:83–88. doi:10.1016/j.mimet.2005.01.010.